## Slide 1: DIA Method Design, Data Acquisition, and Assessment
Jarrett Egertson, Ph.D.

## Slide 2: Part One: Fundamental Method Design

## Slide 3: There is No Universal DIA Method
Duty cycle
Number of Injections
m/z Range Covered
Isolation Width
Resolving Power
AGC Target / Max Inject Time

## Slide 4: Duty Cycle
20 20 m/z-wide windows = 400 m/z
m/z
500
900
~30 seconds
40 10 m/z-wide windows = 400 m/z
m/z
500
900
~30 seconds
10 Hz
15 scans
~7 scans
Duty Cycle: 2 seconds
Duty Cycle: 4 seconds

## Slide 5: Duty Cycle
20 20 m/z-wide windows = 400 m/z
m/z
500
900
~6 seconds
10 Hz
10 40 m/z-wide windows = 400 m/z
m/z
500
900

## Slide 6: Number of Injections
~30 seconds
80 5 m/z-wide windows = 400 m/z
m/z
500
900
m/z
500
900
40 5 m/z-wide windows = 200 m/z
40 5 m/z-wide windows = 200 m/z
~30 seconds

## Slide 7: m/z Range Covered

## Slide 8: Isolation Window Width
Vs.
Vs.
2 m/z
10 m/z
20 m/z
DDA
DIA
Lower precursor selectivity
More peptides co-fragmented
More complex MS/MS spectra
More interference

## Slide 9: Precursor Selectivity
2 m/z
ANFQGAITNR

## Slide 10: Precursor Selectivity
10 m/z
ANFQGAITNR

## Slide 11: Precursor Selectivity
20 m/z
ANFQGAITNR

## Slide 12: Precursor Selectivity
Intensity
4e7
Retention Time (min)
25
26
10 m/z
ANFQGAITNR

## Slide 13: Precursor Selectivity
Intensity
4e7
10 m/z
Retention Time (min)
Intensity
4e7
25
26
20 m/z
ANFQGAITNR
X

X
X



## Slide 14: Precursor Selectivity
SLQDIIAILGMDELSEEDKLTVSR+++
(892.47 m/z)
SLQDIIAILGMDELSEEDKLTVSR+++
(897.8 m/z)
890
900
X


X

## Slide 15: Precursor Selectivity

## Slide 16: Resolving Power

## Slide 17: Precursor Selectivity
SLQDIIAILGMDELSEEDKLTVSR+++
(892.47 m/z)
SLQDIIAILGMDELSEEDKLTVSR+++
(897.8 m/z)
890
900
X


X

## Slide 18: Precursor and Fragment Ion Selectivity
Gallien S, Duriez E., Demeure K, Domon B JPR 2013

## Slide 19: 4 m/z is Key Number for Isolation
Valine
Isoleucine
+ CH2
+2: +7.01 m/z
+3: +4.67 m/z

## Slide 20: Even Better Precursor Selectivity is Useful when using Isotope-Labeled Standards
FDSPESHVGVAWR[+10]
FDSPE SHVGVAWR[+10]
FDSPES HVGVAWR[+10]
FDSPESH VGVAWR[+10]
FDSPESHVGVAWR
FDSPE SHVGVAWR
FDSPES HVGVAWR
FDSPESH VGVAWR
Light Precursor
Heavy Precursor
Heavy y - ions
Light b - ions
Light b - ions
Light y - ions

## Slide 21: Transition Selection for DIA – y-ions only!
FDSPESHVGVAWR[+10]++
748.86 m/z
FDSPESHVGVAWR++
743.86 m/z
5 m/z

## Slide 22: Transition Selection for DIA – y-ions only!
FDSPESHVGVAWR[+10]++
748.86 m/z
FDSPESHVGVAWR++
743.86 m/z
5 m/z
SRM Isolation
0.7 m/z, centered
OK b or y

## Slide 23: Transition Selection for DIA – y-ions only!
FDSPESHVGVAWR[+10]++
748.86 m/z
FDSPESHVGVAWR++
743.86 m/z
5 m/z
DIA Isolation
20 m/z
BAD only y

## Slide 24: Transition Selection for DIA – y-ions only!
FDSPESHVGVAWR[+10]++
748.86 m/z
FDSPESHVGVAWR++
743.86 m/z
5 m/z
DIA Isolation
20 m/z
OK b or y

## Slide 25: Transition Selection for DIA – y-ions only!
FDSPESHVGVAWR[+10]++
748.86 m/z
FDSPESHVGVAWR++
743.86 m/z
5 m/z
DIA Isolation
20 m/z
Still Bad – y only

## Slide 26: AGC Target / Max IT
DDA
MS/MS for peptide identification
For detection – only enough ions to generate peptide-spectrum match
Long fill times may mean slower acquisition rate, less ID’s
DIA
MS/MS for peptide detection and quantification
For quantification – want as many ions as possible
  Precision
  Sensitivity
  Intra-scan dynamic range
Long fill times can slow down duty cycle, hinder quantification

## Slide 27: DIA Parameters Influence Each Other
Duty cycle
Number of Injections
m/z Range Covered
Isolation Width
Resolving Power
AGC Target / Max Inject Time
Duty cycle
Number of Injections
m/z Range Covered
Isolation Width
Resolving Power
AGC Target / Max Inject Time

## Slide 28: Putting Together a DIA Method
Determine Duty Cycle
Choose Isolation Window Width
Determine max IT/ Resolving Power
Determine m/z Range To Cover
Duty cycle
Number of Injections (1)
m/z Range Covered
Isolation Width
Resolving Power
AGC Target / Max IT







## Slide 29: Step 1: Determine Duty Cycle
Required duty cycle based on LC
  At least 7 points across chromatographic peak
Narrow peaks
  Faster duty
  cycle
~15 seconds
15 seconds / 7 points =
2.15 second duty cycle
Determine Duty Cycle

## Slide 30: Step 2: Determine m/z Range to Cover
PRTC Peptides
Determine m/z Range To Cover

## Slide 31: Step 3: Choose Isolation Window Width
For 500 – 900 m/z on
  QE: 15-25 m/z
  QE-HF: 10 – 20 m/z
  Fusion: 10 – 20 m/z
Selectivity
Ion Counts
More important for complex samples

## Slide 32: Determine Required Acquisition Rate
Duty Cycle
2.0 seconds
m/z Range
500 – 900 m/z (400 m/z)
Isolation Width
20 m/z
MS/MS Scans per Duty Cycle:
MS Scans per Duty Cycle: 1 (assume ~75 ms for acquisition)
Required MS/MS Acquisition Rate
= 96 .25 ms / scan = 10.4 Hz
= 20 scans

## Slide 33: Determine IT / Resolving Power (QE-HF)
AGC Target: 1e6

## Slide 34: 40 x 10 m/z Method Underfills (QE-HF)
Max IT: 17 milliseconds
Max IT: 60 milliseconds
Selectivity
Ion Counts
Selectivity
Ion Counts

## Slide 35: A Recommended Starting Point

## Slide 36: A Recommended Starting Point
QE-HF (Increased MS2 Resolving Power)
MS2 Resolving Power: 17,500 -> 30,000
Maximum IT: auto (49 ms) -> 60 ms
Fusion
Similar to QE-HF*
  Orbitrap acquisition is slightly slower
  AGC Target: 2e5

## Slide 37: Part Two: Advanced Concepts

## Slide 38: Advanced Concepts
Optimizing isolation window placement
Isolation uniformity
Resonance CID vs. HCD

## Slide 39: Window Placement
Windows are no longer centered on precursors
696
697
698
699
700
701
702
703
704
705
m/z
0
5
10
15
20
25
30
35
40
45
50
55
60
65
70
75
80
85
90
95
100
Relative Abundance
699.88
700.38
700.89
701.39
696.82
697.32
698.84
699.34
703.41
701.89
703.91
702.86
696.34
704.82

## Slide 40: Peptides Masses Fall in Discrete Bins
1.00045475 m/z
 | Mass | Excess
H | 1.00078 | 0.00078
C | 12 | 0.0
O | 15.9949 | 0.9949
N | 14.0031 | 0.0031
S | 31.9721 | 0.9721

## Slide 41: Window Placement
 | Mass | Excess
H | 1.00078 | 0.00078
C | 12 | 0.0
O | 15.9949 | 0.9949
N | 14.0031 | 0.0031
S | 31.9721 | 0.9721
H
C
N
O

## Slide 42: Window Placement
 | Mass | Excess
H | 1.00078 | 0.00078
C | 12 | 0.0
O | 15.9949 | 0.9949
N | 14.0031 | 0.0031
S | 31.9721 | 0.9721
15.0023
26.0031

## Slide 43: Peptides Masses Fall in Discrete Bins
1.00045475 m/z
 | Mass | Excess
H | 1.00078 | 0.00078
C | 12 | 0.0
O | 15.9949 | 0.9949
N | 14.0031 | 0.0031
S | 31.9721 | 0.9721

## Slide 44: Window Placement

## Slide 45: Window Placement

## Slide 46: Skyline Demonstration
Generating a DIA Isolation List and Using it to Build a QE Method

## Slide 47: Isolation Uniformity
Q-Exactive
Q-Exactive HF

## Slide 48: Fragmentation
Without a targeted precursor
  CE may not be optimal (charge is unknown)

## Slide 49: Fragmentation
Without a targeted precursor
  CE may not be optimal (charge is unknown)

## Slide 50: Fragmentation
Without a targeted precursor
  CE may not be optimal (charge is unknown)

## Slide 51: Fragmentation
Without a targeted precursor
  CE may not be optimal (charge is unknown)

## Slide 52: Fragmentation
Without a targeted precursor
  CE may not be optimal (charge is unknown)

## Slide 53: Comparing reCID to HCD
m/z
400
1000
200 3 m/z-wide windows = 600 m/z
12 seconds total @ 17 Hz scan rate
reCID:
Efficient fragmentation without charge optimization
Generation of b-ion series
HCD
Speed
Preservation of fragment ions within isolated m/z range

## Slide 54: Duty Cycle: reCID: ~16.7 Hz

## Slide 55: Duty Cycle: HCD: ~20 Hz

## Slide 56: Collision Energy
Chart: Unique Peptides (q <= 0.01)
| reCID | HCD
CE30 | 10933
CE35 | 10737 | 7898
Ramped | 7143
Resonance CID May Outperform HCD for a DIA Experiment
C. elegans lysate, database search using SEQUEST

## Slide 57: Part 3: Data Assessment

## Slide 58: Quality Control Overview
QC
QC
QC
QC
Sample
Sample
Sample
Sample
Sample
Sample
QC
Peptide Retention Time Calibration Mixture
# | Peptide Sequence | Mass | Hydrophobicity Factor (HF)
1 | SSAAPPPPPR | 985.5220 | 7.56
2 | GISNEGQNASIK | 1224.6189 | 15.50
3 | HVLTSIGEK | 990.5589 | 15.52
4 | DIPVPKPK | 900.5524 | 17.65
5 | IGDYAGIK | 843.4582 | 19.15
6 | TASEFDSAIAQDK | 1389.6503 | 25.88
7 | SAAGAFGPELSR | 1171.5861 | 25.24
8 | ELGQSGVDTYLQTK | 1545.7766 | 28.37
9 | GLILVGGYGTR | 1114.6374 | 32.18
10 | GILFVGSGVSGGEEGAR | 1600.8084 | 34.50
11 | SFANQPLEVVYSK | 1488.7704 | 34.96
12 | LTILEELR | 995.5890 | 37.30
13 | NGFILDGFPR | 1144.5905 | 40.42
14 | ELASGLSFPVGFK | 1358.7326 | 41.18
15 | LSSEAPALFQFDLK | 1572.8279 | 46.66
PRM

## Slide 59: Skyline QC Demonstration
Generating a QC Method and Analyzing the Data in Skyline

## Slide 60: Quality Control
Targeted-MS2 allows for monitoring of chromatography
Retention time reproducibility is important for DIA (aids peak picking)

## Slide 62: Conclusions
There is no universal DIA method
Try to fill the trap for MS/MS scans
Quality control should monitor chromatography
Determine Duty Cycle
Choose Isolation Window Width
Determine max IT/ Resolving Power
Determine m/z Range To Cover