## Slide 1
Multiplexing Clinical Protein Diagnostics in Dried Blood Spots
Skyline Webinar – March 10, 2015
James Bollinger - MacCoss Lab
Department of Genome Sciences
University of Washington

## Slide 2: Outline
Selection of candidate proteins for the multiplex analysis of DBS via targeted proteomics
The currently employed strategies for the selection of candidate peptides for targeted proteomics
An empirical refinement process for the selection of optimal peptides and their respective MS/MS transitions

## Slide 3: Selection of Protein Targets for DBS Wellness Assay
Adapted from: Vaisar, T. et al., J. Clin. Inves. 2007 117(3); 746–756
CRP
APOB
CP
CFAB
C1R
FIBα/β
FINC
HBB/A1c
LDHB
SYUA
TALDO
ALDOA
KAD1
ZA2G
MDHC
TTHY
VTDB
IgG

## Slide 4
N. Leigh Anderson & Norman G. Anderson Mol Cell Proteomics (2002), 1, 845-867
N. Leigh Anderson. Clin. Chem. (2010), 56, 177-185
Selection of Protein Targets for DBS Wellness Assay
Notes: Reference intervals for 70 protein analytes in plasma. Abundance is plotted on a log scale spanning 12 orders of magnitude. Where only an upper limit is quoted, the lower end of the interval line shows an arrowhead. The classical plasma proteins are clustered to the left (high abundance), the tissue leakage markers (e.g. enzymes and troponins) are clustered in the center, and cytokines are clustered to the right (low abundance). Hemoglobin is included (far left) for comparison. Data were obtained from the Specialty Laboratories publications (154). TPA, tissue plasminogen activator; GCSF, granulocyte colony-stimulating factor; TNF, tumor necrosis factor.

## Slide 5
How Does One Select Optimal Peptides for the Quantitative Analysis of Each Respective Gene Product?
Should one use peptides identified in previous discovery experiments?
Should one use empirical refinement of analytical standards?
Apolipoprotein B (P04114) – 516 kDa, Serum reference range (~0.5-2 g/L)
Walldius, G. et al. Lancet, 2001, 358(9298); 2026-33
Walldius, G. et al. J Intern Med. 2006, 259; 493-519
Notes: Receiver operating characteristic curve: increased discrimination between normal and diseased, i.e. means of distributions separate, using apoB versus LDL-C.
Increased true positive rate (TPR) using apoB instead of LDL-C
  Unique to a given gene product, free of polymorphisms
  Gas-phase properties that favor a sensitive ESI response
  Fragmentation patterns with salient features for accurate detection and quantification
  Stable during/after digestion
  Reproducible chromatographic profiles

## Slide 6: ESPPredictor Score vs. SRM Signal Intensity: Apolipoprotein B100
Mallick, P. & Aebersold, R. et al. Nature Biotechnology(2007) 25 (1): 125-131
Fusaro, V.A. & Carr, S.A. et al. Nature Biotechnology (2009) 27:190-198.
Chart: ESPPREDICTOR sCORE

## Slide 7: Spectrum Counts vs. SRM Signal Intensity: Apolipoprotein B100
Chart: SRM Signal Intensity
# of +2 charge state spectra in 12 independent DDA runs | 7 | 5 | 2 | 4 | 5 | 4 | 7 | 3 | 2 | 8 | 2 | 7 | 5 | 5 | 1 | 2 | 7 | 7 | 4 | 4 | 6 | 3 | 2 | 2 | 6 | 3 | 1 | 3 | 2 | 1 | 10
Prakash, A. et al. J Proteome Res. 2009 8(6): 2733–2739.
Notes: Although these approaches are better than selecting proteotypic peptides at random, our results suggest that current predictor algorithms and spectral counting approaches provide imperfect ranking and identification of optimal proteotypic peptides, potentially limiting the utility of large-scale peptide-synthesis efforts that rely on such approaches as a first-round filter.

## Slide 8: Our Approach…………..

## Slide 9
Workup of Recombinant/Native Protein Standard
http://dnasu.org
Gene Product of Interest
GST
IEAIPQIDK
IEAIPQIDK
+
Expand, Purify, & Sequence cDNA clone
In Vitro Transcription and Translation
Enrichment via Glutathione-Sepharose Resin
~25-500 femtomoles/25uL IVT
< $20.00/Protein
Notes: The 1-Step Human IVT System uses HeLa cell lysates to perform cell-free protein expression. Using a human-based system yields proteins that mirror natural proteins with respect to post-translational modifications and results in proteins that exhibit natural protein function and structure.

## Slide 10
Peptide Filter #1: The y-ion series (y3 to yn-1) for all fully tryptic peptides 7 to 25 amino acids in length (+2 charge state) with all cysteines monitored as carbamidomethyl cysteines.
241
Apolipoprotein B100
Peptides
Round 1 nanoLC MS/MS: SRM Signal Intensity Rank and Fragmentation Fingerprint
In Silico digestion
Export transition list parameters
(m/z Precursors, m/z Products, & Collision Energies1)
1) Maclean, B. et al. Anal Chem. 2010. 82(24):10116-24.
Notes: Before

## Slide 11
Round 1 nanoLC MS/MS:
SRM Signal Intensity Rank
Notes: Feel for peak shape

## Slide 12
Round 1 nanoLC MS/MS:
Fragmentation Fingerprint
R.TGISPLALIK.G [219, 228]
K.ALVEQGFTVPEIK.T [2578, 2590]
K.TTLTAFGFASADLIEIGLEGK.G [674, 694]
R.YEDGTLSLTSTSDLQSGIIK.N [1536, 1555]

## Slide 13
Round 2 nanoLC MS/MS:
iRT Calibration/Peptide Stability Analysis
191
241
Apolipoprotein B100
Peptides
Peptide Filter #2: Edit and annotate all peptides with non-existent or ambiguous chromatograms
Upload curated assay to Panorama for initial build of chromatogram library
www.panorama.org
Export methods with top 4 transitions from remaining peptides (Sensitivity and selectivity arguments) for retention time calibration and stability analysis
Notes: Before

## Slide 14
Round 2 nanoLC MS/MS:
iRT Calibration

## Slide 15
K.DNVFDGLVR.V [4168, 4176]
Round 2 nanoLC MS/MS:
Peptide Stability Analysis
Standard
Standard
72hrs in A/S
R.TGISPLALIK.G [219, 228]
Standard
Standard
72hrs in A/S
K.TTLTAFGFASADLIEIGLEGK.G [674, 694]
Standard
Standard
72hrs in A/S

## Slide 16
Round 3 nanoLC MS/MS:
Validation of Peptides/Transitions in Matrix
191
55
241
Apolipoprotein B100
Peptides
Peptide Filter #3: Edit and annotate all peptides that fail to meet minimum stability requirements.
Export scheduled methods with all transitions from all remaining peptides for the screening of a matrix of interest (DBS)
Update chromatogram library files with iRT values

## Slide 17
Round 3 nanoLC MS/MS:
Validation of Peptides/Transitions in Matrix
K.SPAFTDLHLR.Y [3979, 3988]
R.EFQVPTFTIPK.L [1333, 1343]

## Slide 18
Peptide Filter #4: Edit and annotate all peptides that were not reliably identified in the matrix of interest
Round 4 nanoLC MS/MS:
Digestion Time Course
191
55
241
32
Apolipoprotein B100
Peptides
Combine and export scheduled methods for all remaining peptides for digestion time course analysis

## Slide 19
Round 4 nanoLC MS/MS:
Digestion Time Course
K.ALVEQGFTVPEIK.T [2578, 2590]

## Slide 20
Final Method Refinement
Peptide Filter #5: Edit and annotate all peptides that gave a sub-optimal digestion characteristics
191
57
19
241
32
Apolipoprotein B100
Peptides
Normalization?
APOA1 Global Internal Standard (15N Protein)1
Calibration?
Single Point Calibrator2
Assign concentration with clinical immunoassay
Measure in triplicate in each batch
Hoofnagle, A.N. et al. Clin. Chem. 2012, 58(4); 777-781.
Cox, B. et al. Clin. Chem. 2014, 60(3); 541-8.
Combine and export scheduled methods for all remaining peptides from all target proteins for DBS correlation studies
55 Proteins, 1406 transitions, 281 Peptides

## Slide 21
Summary of Empirical Derivation
with Vitamin D Binding Protein (P02774)
K.VLEPTLK.S [363, 369]
K.HLSLLTTLSNR.V [207, 217]
R.THLPEVFLSK.V [353, 362]
Protein Standard
DBS