## Slide 1: Large Scale DIA With Skyline
Tutorial Webinar #14
With
Brendan MacLean (Principal Developer, Skyline)

## Slide 2: Agenda
Welcome from the Skyline team!
Large Scale DIA with Skyline
  Introduction and overview with Brendan MacLean
  Tutorial with Brendan MacLean
Audience Q&A – submit questions to Google Form:
https://skyline.ms/QA4Skyline.url

## Slide 3: Webinar 2: Jump start DIA… November, 2014

## Slide 4: Chromatography-based Quantification
Hypothesis testing (Verification)
SRM
MS1 chromatogram extraction
Targeted MS/MS (PRM)
Data independent acquisition (DIA/SWATH)
Acquisition | Targeted | Survey
More Selective | PRM | DIA
Less Selective | SRM | MS1
Got HYPOTHESIS??
Exploratory vs. Confirmatory

## Slide 5: Prior Knowledge and Consistency
Relative product ion abundance
  Spectral libraries
Retention time
  iRT libraries
Powerful enough to be used cross-lab / cross experiment
  Measure, store, re-use
More powerful run-to-run
Notes: Confidence in what is being measured

## Slide 6: Relative Product Ion Abundance

## Slide 7: Webinar 7: iRT time prediction… May 2015
Escher, Reiter, et al. Proteomics, 2012

## Slide 8: Defining an iRT Scale
Points on a line (score = time * slope + intercept)
Chart: Measured Time
| 11.286191940307599 | 15.607408523559601 | 16.995498657226602 | 18.1365661621094 | 19.103858947753899 | 19.833576202392599 | 21.436367034912099 | 23.1859245300293 | 25.897150039672901 | 28.410640716552699 | 30.394607543945298
iRT-C18 | -29.2227 | 0 | 9.3871 | 17.1037 | 23.6451 | 28.5799 | 39.4189 | 51.2505 | 69.5855 | 86.5832 | 100

## Slide 9: Building an iRT Library
*
*
*
*
*
*
*
*
*
*
*

## Slide 10: Prior Knowledge Workflow
Chromatogram Extraction
Spectral Library
iRT Library
Product Ion
Targeting
Quantitative
Analysis
DIA Acquisition
Relative Ion Abundance
Normalized RT
mProphet
Peak Picking
Reiter, Rinner, et al.
Nature Methods, 2011

## Slide 11: Prior Knowledge Workflow
DIA Acquisition
Spectronaut
PeakView
Assay Library
Relative Ion Abundance
Normalized RT
Chromatogram Extraction
Product Ion
Targeting
Quantitative
Analysis
mProphet
Peak Picking

## Slide 12: Prior Knowledge Workflow
Chromatogram Library
DIA Acquisition
Relative Ion Abundance
Normalized RT
Chromatogram Extraction
Product Ion
Targeting
Quantitative
Analysis
mProphet
Peak Picking

## Slide 13: Two Phase Experiment
N * Quantitative Samples
1 * Detection Samples
pool
DIA
DDA
DDA
DDA
DDA
DDA
DDA
N *
Quantify Peptides
Detect Peptides
with Search Engine
Build Library
Bruderer, et al. & Reiter, MCP, 2015

## Slide 14: DIA/SWATH Software Comparison
Driving lots of performance improvement
  MaxQuant tutorial MS1 data imports 3x faster
  Comparison pipeline tasks 20x to 200x faster
    Single file import 5 hours on SSD with 3.1
    6 files import in 50 minutes on spinning drive with 3.5
    Adjusting peaks took 2 hours with 3.1
    Adjusting peaks takes 10 seconds with 3.5
Looked bad for Skyline DIA quant
  Turned out to be serious issues with settings
Navarro, et al. & Tenzer
A multicenter study, Nature Biotech. 2016

## Slide 15: SkylineRunner Command-Line Interface
Great for long batch runs
Ready to be adapted to a compute cluster

## Slide 16: Before
Spectronaut 42,439 peptides identified
OpenSWATH 40,387 peptides identified
DIA Umpire 31,256 peptides identified
PeakView 28,424 peptides identified
Skyline 27,121 peptides identified

## Slide 17: Finding the Problem
TPVIVTLK
Spectronaut
Skyline
Wrong!!

## Slide 18: Extraction Width
10,000 rp (default) 30,000 rp

## Slide 19: Extraction Width
60,000 rp ±10 ppm (centroided)

## Slide 20: Extraction from Centroided Spectra

## Slide 21: Extending SkylineRunner
--reintegrate-model-name=<name> | The name of a scoring model to use for the reintegrate operation. The model can either be pre-defined (e.g. using the Edit > Refine > Reintegrate form) or created automatically during this operation by using --reintegrate-create-model.
--reintegrate-create-model | This option will cause a new model to be created, using the mProphet algorithm with all available scores for the results found in the document. (requires --reintegrate-model-name)
--reintegrate-annotate-scoring | Peaks will be annotated with q value and score annotations. (requires --reintegrate-model-name)
--reintegrate-overwrite-peaks | Existing manually integrated peaks will be overwritten with peaks chosen by the reintegration model. (requires --reintegrate-model-name)

## Slide 22: Extending SkylineRunner
--full-scan-precursor-res=<resolving power> | Resolving power of the precursor mass analyzer.
--full-scan-precursor-res-mz=<m/z value> | The m/z value at which the precursor mass analyzer resolving power is specified. (applies only to orbitrap and ft_icr mass analyzers)
--full-scan-product_res=<resolving power> | Resolving power of the product mass analyzer.
--full-scan-precursor-res-mz=<m/z value> | The m/z value at which the product mass analyzer resolving power is specified. (applies only to orbitrap and ft_icr mass analyzers)
--full-scan-rt-filter-tolerance=<minutes> | The number of minutes on either side of the predicted time or MS/MS IDs, i.e. ± minutes. Defaults to.

## Slide 23: Optimal Settings Analysis (RP)
Chart: Resolving Power Settings in 6600 with 64 Variable Windows
| 10 | 30 | 45 | 60 | 90
Peptides | 29495 | 37460 | 38742 | 39074 | 39003
Proteins | 3701 | 4478 | 4588 | 4604 | 4596

## Slide 24: Optimal Settings Analysis (RT)
Chart: RT Window Settings in 6600 with 64 Variable Windows Centroided
| 4 | 6 | 8 | 10 | 12 | 16 | 20 | 24 | 28 | 32
Peptides | 37693 | 40297 | 41146 | 41440 | 41540 | 41718 | 41902 | 42127 | 42288 | 42340
Proteins | 4535 | 4677 | 4718 | 4739 | 4746 | 4764 | 4774 | 4781 | 4796 | 4802

## Slide 25: After
Skyline 42,517 peptides identified
Spectronaut 42,325 peptides identified
OpenSWATH 40,728 peptides identified
DIA Umpire 36,249 peptides identified
SWATH 2.0 35,489 peptides identified

## Slide 26: Multi-File Parallel Loading

## Slide 27: Multi-File Loading Performance
Chart: 12 x WIFF – 24-core, 178 GB, SSD
| 1 | 3 | 6 | 12
processes | 179 | 68.25 | 38.25 | 23.25
threads | 179 | 88.5 | 75.6667 | 82.8333
Chart: 6 x WIFF – 4-core, 12 GB, SSD
| 1 | 2 | 3 | 6
processes | 69 | 41.5 | 35 | 35
threads | 69 | 48 | 44 | 49.5

## Slide 28: Analyzing Data From Navarro Paper
Acquired by L. Gillet

## Slide 29: Blib Document mProphet Model

## Slide 30: Assay Library Document mProphet Model

## Slide 31: Learn More
Webinar #15: TBD (14 recorded webinars)
Weeklong Courses 2017
  Buck Institute, Feb 27 – Mar 3 - Full
  Northeastern University, Boston – May 1-3 – Register now!
  ETH, Zurich, June 26 – 30 – Register now!
  University of Washington, Seattle – July 24-28
Workshops and Conferences 2017
  Skyline User Group Meeting at ASMS, Indianapolis – June 4
  Listings updated in Join Us section of Skyline homepage:
  https://skyline.ms/Skyline.url

## Slide 32: Questions?
Ask any questions at the following form:
  https://skyline.ms/QA4Skyline.url
Take the post-webinar survey:
https://skyline.ms/survey4webinar.url

## Slide 33: Tutorial Webinar #14
This ends this Skyline Tutorial Webinar.
Please give us feedback on the webinar at the following survey:
https://skyline.ms/survey4webinar.url
A recording of today’s meeting will be available shortly at the Skyline website.
We look forward to seeing you at a future Skyline Tutorial Webinar.