## Slide 1
Problem: CE optimization results sometimes produce unexpected peak area graphs or have completely missing data in Skyline. This happens for some or all CE steps for a single ion, and may affect multiple ions for a peptide.
Instrument: Agilent 6495 QQQ, acquisition software version B.07.00
CE optimization transition list is generated in Skyline and imported into Agilent MassHunter and used to acquire data
3 transitions as examples:
  DPENFPFVVLGNK y8 (Good, CE pattern and best CE match between programs)
  QIGDNLIVPGGVK y6 (Not good, CE pattern and best CE is different between programs)
  VSEVKPSYR y3 (Not good, CE pattern and best CE is different between programs)
Product ion masses in data file are sometimes different from the product ion masses in the acquisition method

## Slide 2
  DPENFPFVVLGNK y8 (Good, CE pattern and best CE match between programs)
Acquisition method:

## Slide 3
Peptide Sequence | Replicate Name | Precursor Mz | Precursor Charge | Product Mz | Product Charge | Fragment Ion | Area | Opt Step | CE
DPENFPFVVLGNK | 1 | 742.387191 | 2 | 881.533464 | 1 | y8 | 6296 | -5 | 5
DPENFPFVVLGNK | 1 | 742.387191 | 2 | 881.533464 | 1 | y8 | 15016 | -4 | 9
DPENFPFVVLGNK | 1 | 742.387191 | 2 | 881.533464 | 1 | y8 | 43049 | -3 | 13
DPENFPFVVLGNK | 1 | 742.387191 | 2 | 881.533464 | 1 | y8 | 122526 | -2 | 17
DPENFPFVVLGNK | 1 | 742.387191 | 2 | 881.533464 | 1 | y8 | 242788 | -1 | 21
DPENFPFVVLGNK | 1 | 742.387191 | 2 | 881.533464 | 1 | y8 | 339380 | 0 | 25
DPENFPFVVLGNK | 1 | 742.387191 | 2 | 881.533464 | 1 | y8 | 265251 | 1 | 29
DPENFPFVVLGNK | 1 | 742.387191 | 2 | 881.533464 | 1 | y8 | 183871 | 2 | 33
DPENFPFVVLGNK | 1 | 742.387191 | 2 | 881.533464 | 1 | y8 | 97741 | 3 | 37
DPENFPFVVLGNK | 1 | 742.387191 | 2 | 881.533464 | 1 | y8 | 38828 | 4 | 41
DPENFPFVVLGNK | 1 | 742.387191 | 2 | 881.533464 | 1 | y8 | 10137 | 5 | 45
DPENFPFVVLGNK y8 – Data in Skyline
Normal peak area graph when areas plotted
Largest peak area achieved with CE of 25
Data reported for all CEs for this ion
Data reported for all other ions for this peptide
Exported report for peak area
Best CE with highest peak area: 25

## Slide 4
Chart: DPENFPFVVLGNK y8
| 5 | 9 | 13 | 17 | 21 | 25 | 29 | 33 | 37 | 41 | 45
Area | 9681 | 38688 | 97394 | 185448 | 260386 | 342094 | 243353 | 124884 | 42210 | 14419 | 6726
DPENFPFVVLGNK y8 – Data in MassHunter Qualitative Analysis program
Normal peak area graph when areas plotted
Largest peak area achieved with CE of 25
Data reported for all CEs for this ion
Data reported for all other ions for this peptide
Name | RT | Precursor | m/z (prod.) | Height | Area | CE
DPENFPFVVLGNK.742.387191.881.583464.45 | 37.863 | 742.3872 | 881.5835 | 847 | 9681 | 5
DPENFPFVVLGNK.742.387191.881.573464.41 | 37.808 | 742.3872 | 881.5735 | 3458 | 38688 | 9
DPENFPFVVLGNK.742.387191.881.563464.37 | 37.825 | 742.3872 | 881.5635 | 7818 | 97394 | 13
DPENFPFVVLGNK.742.387191.881.553464.33 | 37.797 | 742.3872 | 881.5535 | 15119 | 185448 | 17
DPENFPFVVLGNK.742.387191.881.543464.29 | 37.834 | 742.3872 | 881.5435 | 20925 | 260386 | 21
DPENFPFVVLGNK.742.387191.881.533464.25 | 37.808 | 742.3872 | 881.5335 | 31763 | 342094 | 25
DPENFPFVVLGNK.742.387191.881.523464.21 | 37.85 | 742.3872 | 881.5235 | 21750 | 243353 | 29
DPENFPFVVLGNK.742.387191.881.513464.17 | 37.851 | 742.3872 | 881.5135 | 10133 | 124884 | 33
DPENFPFVVLGNK.742.387191.881.503464.13 | 37.831 | 742.3872 | 881.5035 | 3890 | 42210 | 37
DPENFPFVVLGNK.742.387191.881.493464.9 | 37.809 | 742.3872 | 881.4935 | 1521 | 14419 | 41
DPENFPFVVLGNK.742.387191.881.483464.5 | 37.788 | 742.3872 | 881.4835 | 543 | 6726 | 45
Exported report for peak area
Best CE with highest peak area: 25

## Slide 5
  QIGDNLIVPGGVK y6 (Not good, CE pattern and best CE is different between programs)
Acquisition method:

## Slide 6
QIGDNLIVPGGVK y6 – Data in Skyline
Strange peak area graph pattern
Largest peak area achieved with CE of 17
No data reported for 2 CEs (Opt Step -4 and Opt Step -5 are completely missing
No data reported at all for 3 other ions for this peptide
Peptide Sequence | Replicate Name | Precursor Mz | Precursor Charge | Product Mz | Product Charge | Fragment Ion | Area | Opt Step | CE
QIGDNLIVPGGVK | 1 | 659.384451 | 2 | 564.359522 | 1 | y6 | 328647 | -3 | 13
QIGDNLIVPGGVK | 1 | 659.384451 | 2 | 564.359522 | 1 | y6 | 339826 | -2 | 17
QIGDNLIVPGGVK | 1 | 659.384451 | 2 | 564.359522 | 1 | y6 | 187954 | -1 | 21
QIGDNLIVPGGVK | 1 | 659.384451 | 2 | 564.359522 | 1 | y6 | 82739 | 0 | 25
QIGDNLIVPGGVK | 1 | 659.384451 | 2 | 564.359522 | 1 | y6 | 23024 | 1 | 29
QIGDNLIVPGGVK | 1 | 659.384451 | 2 | 564.359522 | 1 | y6 | 54033 | 2 | 33
QIGDNLIVPGGVK | 1 | 659.384451 | 2 | 564.359522 | 1 | y6 | 119839 | 3 | 37
QIGDNLIVPGGVK | 1 | 659.384451 | 2 | 564.359522 | 1 | y6 | 186777 | 4 | 41
QIGDNLIVPGGVK | 1 | 659.384451 | 2 | 564.359522 | 1 | y6 | #N/A | 5 | 45
Exported report for peak area
Best CE with highest peak area: 17

## Slide 7
Name | RT | Precursor | m/z (prod.) | Height | Area | CE
QIGDNLIVPGGVK.659.384451.564.309522.5 | 24.79 | 659.3845 | 564.3095 | 1725 | 9955 | 5
QIGDNLIVPGGVK.659.384451.564.319522.9 | 24.786 | 659.3845 | 564.3195 | 4663 | 27854 | 9
QIGDNLIVPGGVK.659.384451.564.329522.13 | 24.799 | 659.4 | 564.35 | 14567 | 79634 | 13
QIGDNLIVPGGVK.659.384451.564.339522.17 | 24.792 | 659.4 | 564.35 | 31632 | 183214 | 17
QIGDNLIVPGGVK.659.384451.564.349522.21 | 24.79 | 659.4 | 564.35 | 62754 | 341240 | 21
QIGDNLIVPGGVK.659.384451.564.359522.25 | 24.785 | 659.4 | 564.35 | 73279 | 319256 | 25
QIGDNLIVPGGVK.659.384451.564.369522.29 | 24.798 | 659.4 | 564.35 | 48075 | 255764 | 29
QIGDNLIVPGGVK.659.384451.564.379522.33 | 24.797 | 659.4 | 564.4 | 31922 | 180496 | 33
QIGDNLIVPGGVK.659.384451.564.389522.37 | 24.785 | 659.4 | 564.4 | 25037 | 120278 | 37
QIGDNLIVPGGVK.659.384451.564.399522.41 | 24.799 | 659.4 | 564.4 | 10590 | 54377 | 41
QIGDNLIVPGGVK.659.384451.564.409522.45 | 24.787 | 659.4 | 564.4 | 4855 | 22916 | 45
QIGDNLIVPGGVK y6 – Data in MassHunter Qualitative Analysis program
Normal peak area graph when areas plotted
Largest peak area achieved with CE of 21
Data reported for all CEs for this ion
Data reported for all other ions for this peptide
Chart: QIGDNLIVPGGVK y6
| 5 | 9 | 13 | 17 | 21 | 25 | 29 | 33 | 37 | 41 | 45
Area | 9955 | 27854 | 79634 | 183214 | 341240 | 319256 | 255764 | 180496 | 120278 | 54377 | 22916
Exported report for peak area
Best CE with highest peak area: 21

## Slide 8
  VSEVKPSYR y3 (Not good, CE pattern and best CE is different between programs)
Acquisition method:

## Slide 9
VSEVKPSYR y3– Data in Skyline
Strange peak area graph pattern
Largest peak area achieved with CE of 25
Peptide Sequence | Replicate Name | Precursor Mz | Precursor Charge | Product Mz | Product Charge | Fragment Ion | Area | Opt Step | CE
VSEVKPSYR | 1 | 358.865435 | 3 | 435.222578 | 1 | y3 | 18786 | -5 | 5
VSEVKPSYR | 1 | 358.865435 | 3 | 435.222578 | 1 | y3 | 22070 | -4 | 9
VSEVKPSYR | 1 | 358.865435 | 3 | 435.222578 | 1 | y3 | 31284 | -3 | 13
VSEVKPSYR | 1 | 358.865435 | 3 | 435.222578 | 1 | y3 | 42665 | -2 | 17
VSEVKPSYR | 1 | 358.865435 | 3 | 435.222578 | 1 | y3 | 85446 | -1 | 21
VSEVKPSYR | 1 | 358.865435 | 3 | 435.222578 | 1 | y3 | 104957 | 0 | 25
VSEVKPSYR | 1 | 358.865435 | 3 | 435.222578 | 1 | y3 | 6 | 1 | 29
VSEVKPSYR | 1 | 358.865435 | 3 | 435.222578 | 1 | y3 | 587 | 2 | 33
VSEVKPSYR | 1 | 358.865435 | 3 | 435.222578 | 1 | y3 | 3109 | 3 | 37
VSEVKPSYR | 1 | 358.865435 | 3 | 435.222578 | 1 | y3 | 8867 | 4 | 41
VSEVKPSYR | 1 | 358.865435 | 3 | 435.222578 | 1 | y3 | 5 | 5 | 45
Exported report for peak area
Best CE with highest peak area: 25
Notes: Peptide Sequence

## Slide 10
VSEVKPSYR y3 – Data in MassHunter Qualitative Analysis program
Normal peak area graph when areas plotted
Largest peak area achieved with CE of 9
Name | RT | Precursor | m/z (prod.) | Height | Area | CE
VSEVKPSYR.358.865435.435.172578.5 | 4.605 | 358.85 | 435.15 | 5996 | 18713 | 5
VSEVKPSYR.358.865435.435.182578.9 | 4.607 | 358.85 | 435.2 | 32037 | 132573 | 9
VSEVKPSYR.358.865435.435.192578.13 | 4.61 | 358.85 | 435.2 | 22761 | 111688 | 13
VSEVKPSYR.358.865435.435.202578.17 | 4.607 | 358.85 | 435.2 | 11338 | 52488 | 17
VSEVKPSYR.358.865435.435.212578.21 | 4.61 | 358.85 | 435.2 | 8706 | 38092 | 21
VSEVKPSYR.358.865435.435.222578.25 | 4.617 | 358.85 | 435.2 | 5637 | 21313 | 25
VSEVKPSYR.358.865435.435.232578.29 | 4.607 | 358.85 | 435.25 | 2162 | 8872 | 29
VSEVKPSYR.358.865435.435.242578.33 | 4.605 | 358.85 | 435.25 | 963 | 3031 | 33
VSEVKPSYR.358.865435.435.252578.37 | 4.616 | 358.85 | 435.25 | 160 | 511 | 37
VSEVKPSYR.358.865435.435.262578.41 |  | 358.85 | 435.25 |  |  | 41
VSEVKPSYR.358.865435.435.272578.45 |  | 358.85 | 435.25 |  |  | 45
Chart: VSEVKPSYR y3
| 5 | 9 | 13 | 17 | 21 | 25 | 29 | 33 | 37 | 41 | 45
Area | 18713 | 132573 | 111688 | 52488 | 38092 | 21313 | 8872 | 3031 | 511
Exported report for peak area
Best CE with highest peak area: 9