Skyline Support Request:

Hello,

I am facing some problems with the Skyline software regarding CoA-thioester (small molecule) analysis. For the LC-MS/MS analysis I use the ABSciex QTRAP 6500+ L-MS/MS system coupled with the HPLC Agilent Infinity 1290 system.

For analysis I use Skyline 24.1.

I attached the Skyline files. In the first file 1 std1-std8 depict the first measurement. The software creates nice peaks, but as they are partially cut, I cannot use them to create a calibration curve.

That’s why I optimized the measuring (scheduled, optimized retention times) method with Skyline and measured my samples again. Results are shown in the first file 1 (std12-std82). Now I have nice complete peaks but the peak for ß-Hydroxybutyryl-CoA is missing.

I checked our old software (Analyst, Sciex), and there the peak is present.

Hydroxybutyryl-CoA has the same precursor mass and transition as Malonyl-CoA but elutes at a different retention time.

Skyline gives out these errors (shown in File 2):

Discarding chromatograms for 'Malonyl-CoA Charge [M+]  Ion [347.100549/347.100549] 347.1000+' because the explicit retention time 3.6 is not between 7.4001 and 9.404883

Where Hydroxybutyryl-CoA will elute..

Even without having explicit retention times in the transition list, the chromatogram seems discarded at the end..

!

?

Why is the Peak for ß-Hydrxybutyryl gone? What do I need to change to get rid of the errors?

Thanks for your help!