## Slide 1: Dilution series
Calibration curve of human EN peptide peptide in dog urine digest
  Peptide Immunoaffinity enrichment
  7 min LC gradient
  PRM analysis
Dilution series | fmol/µl (synthetic EN) | fmol (synthetic EN) on column
A | 0 | 0
B | 0.09765625 | 0.24414063
C | 0.1953125 | 0.48828125
D | 0.390625 | 0.9765625
E | 0.78125 | 1.953125
F | 1.5625 | 3.90625
G | 3.125 | 7.8125
H | 6.25 | 15.625
I | 12.5 | 31.25
J | 25 | 62.5
K | 50 | 125
L | 100 | 250
Peptide
SIS: 10 fmol/µl
12,5 fmol on column

## Slide 2
LOD and LOQ calculation |  |  |  |  | 
Theoretical varying peptide [fmol/µl] | Actual varying ppetide amount [fmol/µl] (skyline CC calculation) | Standard deviation | Actual ratio | Theoretical ratio | S/N
0 | 0.17 | 0.11 |  |  | 0.99
0.09765625 | 0.04 | 0.02 |  |  | 0.26
0.1953125 | 0.45 | 0.23 | 10.2 | 2 | 2.68
0.390625 | 0.40 | 0.15 | 0.9 | 2 | 2.40
0.78125 | 0.74 | 0.03 | 1.9 | 2 | 4.45
1.5625 | 1.32 | 0.26 | 1.8 | 2 | 7.94
3.125 | 2.97 | 0.35 | 2.3 | 2 | 17.87
6.25 | 5.81 | 0.45 | 2.0 | 2 | 34.95
12.5 | 11.17 | 0.60 | 1.9 | 2 | 67.21
25 | 21.64 | 1.82 | 1.9 | 2 | 130.14
50 | 55.27 | 1.26 | 2.6 | 2 | 332.47
100 | 100.12 | 11.33 | 1.8 | 2 | 602.24
Blank and Low Concentration Sample | 
 | Mani et. Al 2012
LOD | 0.301
LOQ | 0.903
LOD = μB +t(1-β) (σB + σS)/√n, where μB is the estimated mean of the blank samples, σBis the standard deviation of the blank samples and σS is the standard deviation of the low concentration samples. The equation assumes that analyte concentration is estimated using the mean of n replicates. Given the LOD, LOQ is estimated as 3 × LOD.

## Slide 3: synthetic EN dilution curve + constant SIS
actual conc. fmol/µl