## Slide 1: More Recently (June 2016)
Bruker reports much higher detections with Spectronaut
  MaxQuant DDA library with 19,500 peptide precursors
Dataset 1
  Spectronaut 10,500 peptide precursors detected
Dataset 2
  Spectronaut 17,500 peptide precursors detected
  Skyline 14,000 peptide precursors detected
25% more!

## Slide 2: Mass Error Issue
First clue histograms
Mean = 8.0
SD = 4.1
Mass Error (PPM)
Dataset 2
Frequency
Mean = 0.47
SD = 2.3
Mass Error (PPM)
Dataset 1
Frequency

## Slide 3: Spectronaut Mass Error Correction
Dataset 1
Dataset 2

## Slide 4: Mass Errors from Skyline Reports (R plots)
Dataset 1
Dataset 2

## Slide 5: Mass Errors from Skyline Reports (R plots)
Dataset 1
Dataset 2
Read-Time Calibration Off

## Slide 6: Skyline Mass Error Graph – Data Set 2

## Slide 7: Bruker Re-calibrated m/z – Data Set 2

## Slide 8: After Bruker Calibrated Dataset 2 (July 2016)
Dataset 1
  Skyline 13,800 peptide precursors detected
  Spectronaut 10,500 peptide precursors detected
Dataset 2
  Skyline 16,700 peptide precursors detected
  Spectronaut 17,500 peptide precursors detected